## Slide 1
The INFLUX Experiment
A 12 station surface-based measurement network monitoring CO2, CO, and CH4 has been deployed in and around the Indianapolis, IN metropolitan area as part of the Indianapolis Flux Experiment (INFLUX). Measurements began in 2010 with network installation completed in 2013. Observations are made at heights ranging from 39 m to 130 m using communication towers.
Current State of Tower Measurements:
Cavity ring-down spectrometers (Picarro, Inc.) used at all 12 sites
  Each site has at least one calibrated reference tank sampled daily.
2 sites measure CO2, CO, and CH4
3 sites measure CO2 and CO
7 sites measure CO2 and CH4
  4 CO2-only replaced with CO2 and CH4 in Fall 2014
Flasks packages at 5 sites provide
  One-hour integrated air samples
This poster will discuss:
Instrument calibration procedures
Round-robin test
Flask – In Situ data comparison
Total data availability
CO2, CO, and CH4 Surface In-Situ Measurement Network Uncertainty Analysis and Calibration Procedures in Support of the INFLUX Experiment
Scott Richardson1, Natasha Miles1, Thomas Lauvaux2, Kenneth Davis1,
Jocelyn Turnbull3,4, Anna Karion3,Colm Sweeney3
1Penn State University, 2Jet Propulsion Lab, 3NOAA/University of Colorado, 4GNS Science
Instrument Calibration Procedure
Prior to deployment each CRDS calibrated using 4 NOAA primary standards
  CO2 : 356 – 435 ppm, X2007 scale
  CO: 130 – 303 ppb, WMO CO X2014 scale
  CH4 : 1700 – 2400 ppb, NOAA04 scale
Round Robin Calibration
3 tanks tested 1-5 times at each site, treated as unknowns.
  CO2 : 395.76 ppm / 433.17 ppm / 451.17 ppm
  CO: 143.56 ppb / 254.49 ppb / 305.67 ppb
  CH4 : 1865.29 ppb / 2117.14 ppb / 2350.98 ppb
Sites 01 and 09 calibrated immediately prior to round robin
Site 08 calibrated August 2013
Sites 03, 10, 11, and 13 calibrated July 2011 at NIST
Sites 02, 04, 05, 06, and 07 calibrated following round robin (“old” calibrations”)
Flask Comparison
Flasks filled to provide a one-hour integrated air sample (Turnbull et al., 2012)
Hourly average in situ values compared to flask CO2, CO and CH4 values
The mean offset between in situ and flask measurements is determined for all measurements
Then subset of measurements with low variability in the in situ measurements within the hour
  CO2 < 0.5 ppm, CO < 15 ppb, CH4 < 7 ppb
The low variability subset should provide a better comparison.
CO2 mean difference of 0.05 ± 0.22 ppm (is consistent across all towers and instruments).
CH4 Biases within WMO 2 ppb tolerance
CO Biases within WMO 5 ppb (urban) tolerance
Measurement Network Plans
Maintain 12 measurement sites
Conduct round robing testing with full re-calibration using 4 tanks every 1-2 years
Move one site to NW of city to investigate background site effect on inversion results
CONTACT: Scott Richardson, (814) 574-3232
srichardson@psu.edu
5th NACP Principal Investigators Meeting January 2015, Session: Urban Carbon
All data mean diff: 0.13 +- 0.54 ppm
Low variability mean diff: 0.05 +- 0.22 ppm
All data mean diff: 0.8 +- 4.9ppb
Low variability mean diff: 0.3 +- 4.1 ppb
All data mean diff: -3.4 +- 5.6 ppb
Low variability mean diff: -3.6 +- 4.2 ppb
Round Robin Results
Very recent calibration at sites 01, 08, and 09
  Ave site error: -0.08 to +0.05 ppm CO2
~2 year old calibration at sites 03, 10, 11, and 13
    Ave site error: -0.15 to -0.08 ppm CO2
> 2 year old calibration at site 02, 04, 05, 06, and 07
  Ave site error: -0.25 to +0.1 ppm CO2
Time since last cal more important than age of instrument
CO2 only first generation systems (circled)
  -0.18 to 0.1 ppm average
CO2 Round Robin
CO Round Robin
CH4 Round Robin